## Slide 1
Raman calibration steps:
i) Obtain pixel number for each peak of the cyclohexane spectrum in the low and high Raman shift ranges.
ii) Use a standard reference which gives the actual Raman shift (cm-1) for each peak in the cyclohexane, tolune, and acetonitrile spectrum, for example, cyclohexane can be found in following database.
https://www.chem.ualberta.ca/~mccreery/ramanmaterials.html
iii) Use (i) and (ii) to plot pixel (x-axis) versus Raman shift (y axis) for each peak and then fit a quadratic curve.
iv) In a column of your Origin table, use the resulting quadratic formula to calculate the calibrated Raman shift corresponding to each pixel number.

## Slide 2
Pixel | All (wavenumber) | Cyclohexane | Toluene | Acetonitrile
152 | 384.1 | 384.1 |  | 
164 | 426.4 | 426.4 |  | 
283 | 801.3 | 801.3 |  | 
359 | 1028.3 | 1028.3 |  | 
406 | 1157.6 | 1157.6 |  | 
445 | 1266.4 | 1266.4 |  | 
513 | 1444.4 | 1444.4 |  | 
194 | 521.7 |  | 521.7 | 
278 | 786.6 |  | 786.6 | 
351 | 1003.6 |  | 1003.6 | 
361 | 1030.6 |  | 1030.6 | 
425 | 1211.4 |  | 1211.4 | 
578 | 1605.1 |  | 1605.1 | 
151 | 380 |  |  | 380
323 | 921 |  |  | 921
487 | 1380 |  |  | 1380
514 | 1450 |  |  | 1450
871 | 2256 |  |  | 2256
Low energy region
Chart: Low energy region
| 152 | 164 | 283 | 359 | 406 | 445 | 513 | 194 | 278 | 351 | 361 | 425 | 578 | 151 | 323 | 487 | 514 | 871
All | 384.1 | 426.4 | 801.3 | 1028.3 | 1157.6 | 1266.4 | 1444.4 | 521.7 | 786.6 | 1003.6 | 1030.6 | 1211.4 | 1605.1 | 380 | 921 | 1380 | 1450 | 2256
Cyclohexane | 384.1 | 426.4 | 801.3 | 1028.3 | 1157.6 | 1266.4 | 1444.4
Toluene | 521.7 | 786.6 | 1003.6 | 1030.6 | 1211.4 | 1605.1
Acetonitrile | 380 | 921 | 1380 | 1450 | 2256

## Slide 3
Cyclohexane : Standard 1
https://www.oldcitypublishing.com/journals/lie-home/lie-issue-contents/lie-volume-32-number-5-6-2015/lie-32-5-6-p-295-302/
Lasers in Eng., Vol. 32, pp. 295–302
Standard peaks were resolved during the experiment.

## Slide 4
Toluene : Standard 2
Lasers in Eng., Vol. 32, pp. 295–302
https://www.oldcitypublishing.com/journals/lie-home/lie-issue-contents/lie-volume-32-number-5-6-2015/lie-32-5-6-p-295-302/
Standard peaks were resolved during the experiment.

## Slide 5
Acetonitrile : Standard 3
THE JOURNAL OF CHEMICAL PHYSICS 124, 244312 2006
Standard peaks were resolved during the experiment.
https://doi.org/10.1063/1.2208350

## Slide 6
High energy region
 | All | Cyclohexane | Toluene | Acetonitrile | 
166 | 2664.4 | 2664.4 |  |  | 
233 | 2852.9 | 2852.9 |  |  | 
258 | 2923.8 | 2923.8 |  |  | 
263 | 2938.3 | 2938.3 |  |  | 
308 | 3057.1 |  | 3057.1 |  | 
266 | 2946 |  |  | 2946 | 
288 | 3004 |  |  | 3004 | 
Chart: High energy region
| 166 | 233 | 258 | 263 | 308 | 266 | 288
All | 2664.4 | 2852.9 | 2923.8 | 2938.3 | 3057.1 | 2946 | 3004
Cyclohexane | 2664.4 | 2852.9 | 2923.8 | 2938.3
Toluene | 3057.1
Acetonitrile | 2946 | 3004

## Slide 7
Cyclohexane : Standard 1
After baseline subtraction
https://www.oldcitypublishing.com/journals/lie-home/lie-issue-contents/lie-volume-32-number-5-6-2015/lie-32-5-6-p-295-302/
Lasers in Eng., Vol. 32, pp. 295–302

## Slide 8
Toluene : Standard 2
https://www.oldcitypublishing.com/journals/lie-home/lie-issue-contents/lie-volume-32-number-5-6-2015/lie-32-5-6-p-295-302/
Lasers in Eng., Vol. 32, pp. 295–302

## Slide 9
Acetonitrile : Standard 3
https://doi.org/10.1063/1.2208350
THE JOURNAL OF CHEMICAL PHYSICS 124, 244312 2006